## BOM
Item | Part | Spec to order | Make / Buy | Qty per kit | Unit cost ₹ (prototype) | Line cost ₹ | Price basis | Supplier candidates
1 | Lid with fill chute and window | 304 lid Ø168 × 0.8–1.2 pressed or spun, Ø16 chute, funnel, Ø20 glass window | Make | 1 | 2500 | 2500 | Estimate, one-off fabrication | SS fabricators, Peenya / Jigani (Bangalore)
2 | Lid interlock + lock | 12 mm reed switch + magnet; 24 V mini push solenoid | Buy | 1 | 420 | 420 | Reed ₹28 (Robocraze) + magnet + solenoid estimate | Robocraze; Robu; local electronics
3 | Bowl, spout and flow pins | 304 sheet 0.8 mm rolled Ø150 × 226, welded floor, boss for the seal, Ø10 spout, 2 × 316 Ø1.6 pins | Make | 1 | 3000 | 3000 | Estimate, one-off fabrication | SS fabricators, Peenya / Jigani
4 | Sleeve clamp ring | 316 Ø66 × 4, ID 52, knurled | Make | 1 | 900 | 900 | Estimate, CNC one-off | CNC job shops, Peenya
5 | Sleeve O-ring | Silicone 51 × 1.6, food grade (from a kit) | Buy | 1 | 60 | 60 | Share of a ₹600 silicone O-ring kit | Cart2India / IndiaMART O-ring kits
6 | Filter sleeve, two layers (consumable) | Whatman 602 h, 240 mm circles (cat. 10312620): 4 per sleeve | Buy | 1 | 543.4458 | 543.4458 | Formula: pack price × rate × import ÷ 100 × circles per sleeve | Cytiva/Whatman via Indian lab distributors (Biomall, Chemopharm)
7 | Mesh wrap, 50 µm | 316L 300-mesh twill, 1 sq ft (uses 0.037 m²) | Buy | 1 | 1000 | 1000 | Estimate; IndiaMART 316 filter cloth lists ₹50–450/sq ft | Banaraswala Wire Mesh; Ambani Metals; Satish Wire
8 | Basket with welded vanes | 316L perforated 0.6 mm Ø2/3.1 mm hex, 300 × 300 piece + 316 floor/hub + 4 vanes; roll, seam and TIG | Make | 1 | 4200 | 4200 | Estimate: sheet ₹1,500 + CNC ₹800 + laser ₹400 + welding ₹1,500 | Perforated sheet: Adarsh Metal Perforators, Excel Metal (Mumbai); welding: Peenya
9 | Sleeve floor gasket | Silicone sheet 300 × 300 × 1.5, food grade; cut rings | Buy | 1 | 350 | 350 | Estimate | IndiaMART silicone sheet suppliers
10 | Drive coupler | 316 Ø18 × 6, Ø5 bore, M3 grub | Make | 1 | 700 | 700 | Estimate, CNC one-off; 316 bar ≈ ₹240/kg | CNC job shops, Peenya
11 | Shaft lip seal | FKM 5 × 12 × 4 | Buy | 1 | 250 | 250 | Estimate | Seal stockists (Moglix, local bearing shops)
12 | Base housing | PETG print ≈ 250 g, 0.6 mm walls | Make | 1 | 1500 | 1500 | Estimate, Bangalore print services ≈ ₹5–7/g | Local FDM print services
13 | Feet ×4 | Silicone bumpers Ø14 | Buy | 1 | 120 | 120 | Estimate | Hardware stores
14 | Motor 42BLS02 | 24 V, 4,000 rpm, 0.125 N·m, 52.5 W, 61 mm; confirm hall leads | Buy | 1 | 2080 | 2080 | robu.in listing, Oct 2026 | Robu.in; Aerokart; Robodo (JK42BLS04, ₹3,200, if more torque is needed)
15 | Controller | ESP32 DevKit V1 ₹405 + ZS-X11H 6–60 V hall BLDC driver ₹1,299 + knob, LED, toggle, perfboard ₹400 | Buy | 1 | 2104 | 2104 | techiesms.com + probots.co.in listings + estimate | Probots; Techiesms; Robocraze
16 | 24 V 5 A supply | SMPS 100–240 V in, 24 V 5 A out, plus DC jack and cable | Buy | 1 | 807 | 807 | robocraze.com listing ₹707 + ₹100 cable | Robocraze; Probots (Mean Well LRS series)
17 | Fasteners | A4 stainless M3/M4 screws, grub screws | Buy | 1 | 300 | 300 | Estimate | Local
18 | Electropolish + passivation | All 304/316 wet parts, ASTM A967 | Buy | 1 | 1500 | 1500 | Estimate per lot | SS finishing shops, Peenya
| Prototype kit, all lines (₹) |  |  |  |  | 22334.4458
| Kits to build |  |  |  |  | 2
| Prototype hardware budget (₹) |  |  |  |  | 44668.8915
| Sleeve stock: one pack of 100 circles = 25 sleeves (₹) |  |  |  |  | 13586.144
| Total with sleeve stock (₹) |  |  |  |  | 58255.0355
| PD-298 prototyping cap (₹) |  |  |  |  | 50000
| Within cap? |  |  |  |  | no
Blue on yellow: estimate or choice to replace with an RFQ quote. Green: pulled from the Assumptions tab. Black: formula.
Note on F2: One-off prototype estimates for Bangalore job shops. Replace every yellow cell with a quote.

## Teardown
Item | Part | Function | Material | Process | Mass g (drawn) | Mates with | Change in v1.1 (TRIZ / FMEA ref.)
1 | Lid with fill chute and window | Closes the bowl; Ø16 chute and funnel dose the batch; glass window to watch the drain | 304, 1.2 mm; borosilicate window | Spin/press lid, roll + TIG chute, TIG funnel | 474.6 | 3 bowl rim; 2 interlock tab | Lid held locked while the basket turns (FMEA 15)
2 | Lid interlock + lock | Reed switch proves the lid is shut; solenoid pin holds it until speed = 0 | Bought: reed switch, magnet, 24 V mini solenoid | Buy; mount on base deck | 25 | 1 lid tab; 15 controller | Lock added (FMEA 15, PC4)
3 | Bowl, spout and flow pins | Catches the water off the cage; spout to the jug; two pins sense water in the spout | 304, 1.0 mm (drawn), 0.6–0.8 mm in prototype | Roll + seam weld, weld floor and boss, roll rim; tube spout; 316 pins | 1310.6 | 12 base; 11 seal in the boss; 1 lid | Flow pins added: batch-end and end-of-life signal (principle 23 Feedback; FMEA 10, 12, 14; req. 4.2)
4 | Sleeve clamp ring | Clamps the sleeve top onto the cage rim | 316 | CNC turn, knurl | 39.3 | 8 basket; 5 O-ring; 6 sleeve | —
5 | Sleeve O-ring | Seals the sleeve top edge | Silicone 60 ShA, food grade | Buy | 0.4 | 4 ring; 6 sleeve | —
6 | Filter sleeve, two layers | The filter and the sample; inner layer catches most, outer layer catches what passes and covers inner seams | Cellulose, Grade 602 h class, 2 layers | Cut 4 strips from 240 mm circles; 20 mm overlaps, seams 90° apart | 6 | 8 cage (via 7 mesh); 9 gasket; 5 O-ring | Second layer (TC1, PC1); 20 mm overlaps pressed shut by the spin (principle 25)
7 | Mesh wrap, 50 µm | Holds the wet paper over the Ø2 holes | 316L 300-mesh twill | Cut, wrap, spot-weld seam | 9.1 | 8 cage | —
8 | Basket with welded vanes | Spins the batch; vanes drag the water to ≈ 85 % of basket speed | 316L perforated 0.6 mm, Ø2 at 40 % open; 4 vanes 1 mm; floor + hub | Laser cut, roll + seam weld, TIG vanes to floor and tie ring | 245.2 | 10 coupler; 6 sleeve; 4 ring | Vanes welded in: one part fewer to remove (TC3, principle 10; FMEA 9, 17); vane tips at r 28 leave room to slide the sleeve in
9 | Sleeve floor gasket | Seals the sleeve's bottom edge to the floor | Silicone sheet 1.5 mm, food grade | Die/knife cut | 0.6 | 8 basket floor; 6 sleeve | New: closes the bottom bypass (FMEA 2)
10 | Drive coupler | Joins the basket to the Ø5 motor shaft | 316 | CNC turn, M3 grub screw | 11.2 | 14 shaft; 8 basket | —
11 | Shaft lip seal | Keeps water out of the dry base | FKM, 5 × 12 × 4 | Buy | 0.6 | 3 bowl boss; 14 shaft | Boss raised 14 mm above the floor; drip path (FMEA 7)
12 | Base housing | Holds motor, controller, controls; dry side | PETG (prototype) / PC-ABS (product) | FDM print / injection mould | 205.5 | 3 bowl; 13 feet; 14, 15 | —
13 | Feet ×4 | Grip and isolate vibration | Silicone | Buy | 3 | 12 base | —
14 | Motor 42BLS02 | 24 V BLDC, 4,000 rpm rated, 0.125 N·m, hall sensors give the true speed | Bought | Buy | 430 | 12 base via mount plate; 10 coupler | Closed-loop speed from the hall signals (FMEA 6)
15 | Controller: ESP32 + BLDC driver | Runs the speed profile (≤ 5× face velocity, up to 2,500 rpm), interlock, lock, flow pins | Bought: ESP32 DevKit, ZS-X11H driver | Buy; wire on perfboard | 75 | 14 motor; 2 lock; 3 pins; 16 supply | Speed profile (TC2, PC2); end-of-life alert (req. 4.2)
16 | 24 V 5 A supply | Power | Bought SMPS | Buy | 0 | 15 controller | External; not drawn
| Total mass as drawn (g) |  |  |  | 2836.1
| Spinning mass: items 4–10 (g) |  |  |  | 311.8
Masses from the teardown geometry (Pappus volumes × density). The drawn bowl and lid are 1.0–1.2 mm; at 0.6–0.8 mm the unit is about 0.6 kg lighter.

## Assumptions
Assumption | Value | Unit | Source / note
USD → INR exchange rate | 88 | ₹/$ | Assumed for Oct 2026; edit to the day's rate
Import, shipping and duty on imported lab paper | 30% | fraction | Assumed; Indian lab distributors quote landed prices on login
Grade 602 h, 240 mm circles, pack of 100 | 118.76 | $/pack | scientificfilters.com list (cat. 10312620), Oct 2026
602 h circles per two-layer sleeve | 4 | circles | A 240 mm circle yields a 158 × 180 mm strip; 2 strips per layer, 2 layers
Prototype kits to build | 2 | kits | One to test, one to break
PD-298 prototyping cap | 50000 | ₹ | Department cap used for earlier MPx concepts; edit if different
Sleeve life, tap water (simulated) | 228 | L | v11.mjs: litres until the batch cycle falls below 15 L/h; cake resistance assumed
Sleeve life, borewell (simulated) | 57 | L | same model, 2 mg/L solids
Sleeve life, turbid surface (simulated) | 11.4 | L | same model, 10 mg/L solids
Blue text on yellow = input you can edit; every other number in the workbook is a formula or comes from the drawings.

## Running cost
Water | Sleeve life (L, simulated) | Sleeve cost ₹ | Cost per litre ₹ | Sleeves per year at 3 L/day | Sleeve cost per year ₹
Tap | 228 | 543.4458 | 2.3835 | 4.8026 | 2609.9698
Borewell | 57 | 543.4458 | 9.5341 | 19.2105 | 10439.8791
Turbid surface | 11.4 | 543.4458 | 47.6707 | 96.0526 | 52199.3954
Sleeve life = litres until the batch cycle falls below 15 L/h (cake on the sleeve, α = 10¹³ m/kg assumed). Bench test 4 replaces it.

## Suppliers
Item | Supplier / listing | Location | What they list | Ask in the RFQ | Source
14 | Robu.in — 42BLS02 | Pune (online) | 24 V BLDC 4,000 rpm, 0.125 N·m, 52.5 W, 61 mm, 430 g, ₹2,080 | Confirm hall-sensor leads and 5 mm shaft | https://st.robu.in/product/42bls02-24v-rpm4000-0-125-n-cm-61mm
14 alt | Robodo — JK42BLS04 | online | 24 V, 4,000 rpm, 0.25 N·m, 105 W, Ø5 shaft, hall sensors, 100 mm, ₹3,200 | If spin-up torque is short | https://robodo.in/products/jk42bls04-brushless-motor-105-watt-bldc
15 | Probots — ZS-X11H | Bangalore (online) | 6–60 V, 400 W hall BLDC driver, PWM / 0–5 V speed input, ₹1,299 | Speed-pulse output for closed loop | https://probots.co.in/zs-x11h-400w-bldc-3-phase-brushless-motor-controller-6v-60v-pwm-hall-driver-module.html
15 | Techiesms — ESP32 DevKit V1 | online | ESP32-WROOM-32 board, ₹405 | — | https://techiesms.com/product/esp32-development-board-doit-devkit-v1/
16 | Robocraze — 24 V 5 A SMPS | Bangalore (online) | 100–240 V in, 24 V 5 A, IP20 metal case, ₹707 | Enclosed desktop adapter for the product | https://robocraze.com/products/24v-5a-smps-power-supply
2 | Robocraze — 12 mm reed switch | online | NO reed, 1.2 A, 24 V DC, ₹28 | Pair with a 24 V mini solenoid for the lock | https://robocraze.com/products/12mm-reed-switch
6 | Whatman 602 h circles | Cytiva via distributors | 602 h, 100 per pack: 125 mm $37.93, 150 mm $55.06, 185 mm $80.87, 240 mm $134.29 list ($118.76 sale) | Ask Cytiva India for 602 h in sheets or rolls (no 602 h sheets were found listed) | https://scientificfilters.com/whatman/6-602h
6 alt | Whatman Grade 42, 46 × 57 cm sheets | Chemopharm (India) | Cat. 1442-917, 100 sheets; 8–12 weeks | Slower and coarser (2.5 µm) than 602 h; fallback only | https://store.chemopharm.com/no-42-filter-sheets-46x57cm-100-wh-ma-bo-1442917
7 | IndiaMART filter wire cloth | Coimbatore / Mumbai | 316 filter cloth ₹50–450 per sq ft (Banaraswala 400-mesh ₹50/sq ft; Ambani Metals 150-mesh twill ₹450/sq ft) | 300-mesh 316L twill, 1 sq ft, mill certificate | https://m.indiamart.com/impcat/filter-wire-cloth.html
8 | Perforated sheet (L&T SuFin listings) | Mumbai | SS 316 perforated sheets 1,250 × 2,500 (Adarsh Metal Perforators) | 0.6 mm 316L, Ø2 holes, 3.1 mm 60° pitch, 300 × 300 cut piece | https://lntsufin.com/product/adarsh-metal-perforators-2-mm-stainless-steel-perforated-sheet-for-screening-and-fencing-1250-x-2500-mm-ss-316-astm-a240/16120-466
4, 8, 10 | SS 316 round bar | Raigad (L&T SuFin) | Category RFQ prices ₹226–250/kg (Jul–Sep 2026) | Ø20 and Ø70 offcuts | https://lntsufin.com/product/sps-ss-316-20-mm-stainless-steel-round-bars-6-m/17330-457
— | Carson MicroFlip MP-250 (already owned) | Tanotis (India) | 100–250× pocket microscope, ₹3,390 incl. GST and import | Spare unit for the second kit | https://tanotis.com/products/carson-microflip-mp-250-100-250x-pocket-microscope

## FMEA
# | Part / function | Failure mode | Effect on user | Cause | Control (v1) | O | S | D | RN | Critical? | Action in v1.1 | O new | S new | D new | RN new | Critical? | Req. | New in v1.1
1 | Sleeve paper | Capture falls at high face velocity | 1–2 µm particles pass | Rated at 1/7 of the working velocity | None | 6 | 9 | 9 | 486 | critical | Speed profile ≤ 5× face velocity; second layer; bead test 5 | 2 | 9 | 6 | 108 |  | 6.1
2 | Sleeve seam and ends | Water bypasses the paper | Unfiltered particles reach the jug | Seam or ends unsealed (rig K 3.2× datasheet) | Clamp ring + O-ring | 6 | 9 | 8 | 432 | critical | 20 mm overlaps pressed by spin; seams 90° apart in 2 layers; floor gasket; dye test | 2 | 9 | 5 | 90 |  | 6.1
3 | Detection | Cellulose fibres or stain read as plastic | False positives | Nile red uptake by cellulose | None | 8 | 7 | 7 | 392 | critical | Blank sleeve per batch of sleeves; fibre-shape exclusion; reference counts | 4 | 7 | 4 | 112 |  | 6.2
4 | Spent sleeve | Binned loose or rinsed | Microplastics back to the environment | No disposal route | None | 7 | 6 | 8 | 336 | critical | Sealed sleeve pouch in the box; return-to-lab route for the study | 3 | 6 | 6 | 108 |  | 10.1
5 | Light polymers | PP / PE miss the paper or land unevenly | Sample under-counts them | Inward drift beats through-flow above ≈ 20 µm | None | 6 | 6 | 8 | 288 | critical | Slow stop; floor coupon counted; PP bead test | 4 | 6 | 5 | 120 |  | 1.2
6 | Speed control | Speed drifts out of the band | Low flow or uncertain capture | Open-loop PWM, no feedback | None | 6 | 6 | 8 | 288 | critical | Hall-sensor closed loop on the ESP32; speed logged | 2 | 6 | 2 | 24 |  | 2.1
7 | Lip seal | Leaks into the dry base | Water on the motor and board | Wear at speed, dry running | Lip seal | 4 | 7 | 7 | 196 | critical | Boss 14 mm above the floor; drip path out of the base; board on stand-offs | 3 | 5 | 4 | 60 |  | 3.5
8 | Wet sleeve | Tears or creases at 55–150 g | Bypass; fibres in the water | Wet strength over Ø2 holes | Mesh support | 4 | 8 | 6 | 192 | critical | Two layers; inspect on lift-out; 20-batch test 7 | 3 | 8 | 4 | 96 |  | 6.1
9 | Sleeve change | Sample touched or dropped | Contaminated sample, spilled catch | Four-step removal, wet hands | None | 5 | 6 | 6 | 180 | critical | Three steps (vanes welded in); tweezer tabs; gloves; pouch | 3 | 6 | 4 | 72 |  | 10.2
10 | Speed profile | Starts fast with a full basket | Face velocity ≈ 20× at the start | Firmware fault or wrong batch size | — (new in v1.1) | 3 | 7 | 7 | 147 | critical | Start capped at 1,300 rpm until the flow pins see water; batch size fixed by the funnel | 2 | 7 | 4 | 56 |  | 6.1 | yes
11 | Detection | Too few particles per field | Count meaningless | Sparse sample on a large sleeve | None | 9 | 6 | 2 | 108 |  | Coupons after ≥ 20 L on one sleeve | 4 | 6 | 2 | 48 |  | 6.2
12 | Silt | Sleeve blinds early | Flow under 15 L/h | Borewell / turbid water | None | 7 | 5 | 3 | 105 |  | Life 228 / 57 / 11 L; alert from drain time | 4 | 5 | 2 | 40 |  | 9.1
13 | Cellulose sleeve | Sheds fibres into the water | Particles in the output | Paper rubbing on the mesh | None | 5 | 3 | 7 | 105 |  | 200 mL first flush to waste | 3 | 3 | 5 | 45 |  | 3.2
14 | Flow pins | Pins foul with lime | Wrong batch-end signal | Scale on 316 pins | — (new in v1.1) | 5 | 4 | 5 | 100 |  | Pins wiped at each sleeve change; timed fallback profile | 3 | 4 | 4 | 48 |  | 4.2 | yes
15 | Lid | Opened while spinning | Hand near a basket at 7.9 m/s | Interlock bypassed | Reed switch | 3 | 8 | 4 | 96 |  | Solenoid lock until hall speed = 0 | 2 | 8 | 2 | 32 |  | safety
16 | Cost | Price over ₹6,000 | Not affordable | Steel parts, BLDC drive | None | 8 | 6 | 2 | 96 |  | Rolled parts, printed base; production costing next | 6 | 6 | 2 | 72 |  | 11.1
17 | Vanes | Water lags (< 50 % of basket) | Flow under 15 L/h | Too few or short vanes | 4-vane insert | 5 | 6 | 3 | 90 |  | Welded vanes; bench test 1; profile adapts to drain time | 3 | 6 | 2 | 36 |  | 2.1
18 | Basket | Unbalance from an uneven wet sleeve | Vibration, noise | Uneven cake or sleeve | None | 6 | 4 | 2 | 48 |  | Pre-spin seats the sleeve; feet; 5 g → 10 N at 2,500 rpm checked | 4 | 5 | 2 | 40 |  | 8.1
19 | Motor / driver | Stalls or overheats | No flow | Start torque, no current limit | None | 3 | 5 | 3 | 45 |  | Soft start; driver current limit | 2 | 5 | 3 | 30 |  | 2.1
20 | Bowl | Spray over the rim | Wet bench | Water leaves the cage at ≈ 7 m/s | Skirt | 5 | 3 | 2 | 30 |  | Skirt kept; lid closes the rim | 3 | 3 | 2 | 18 |  | —
| Critical modes | [7 empty] | 10 | [5 empty] | 0
| Sum of RN | [7 empty] | 3760 | [5 empty] | 1255
O, S, D 1–10 (Pahl & Beitz §10.4). Draft ratings in blue on yellow: change them in the team session and RN and the counts update.

## TRIZ
ID | Contradiction | Improving parameter | Worsening parameter | Matrix principles | Principle used | Idea | Verdict
TC1 | Capture vs flow | #30 Object-affected harmful factors | #39 Productivity | 22, 35, 13, 24 | 24 Intermediary | a second sleeve layer catches what the first misses and covers its seams | adopted
TC1 | Capture vs flow | #30 Object-affected harmful factors | #39 Productivity | 22, 35, 13, 24 | 10 Preliminary action | flush a new sleeve with 200 mL to waste before use | adopted
TC1 | Capture vs flow | #30 Object-affected harmful factors | #39 Productivity | 22, 35, 13, 24 | 22 Blessing in disguise | silt cake as an extra depth filter | not relied on
TC1 | Capture vs flow | #30 Object-affected harmful factors | #39 Productivity | 22, 35, 13, 24 | 13 The other way round | make particles bigger with a coagulant | rejected: adds a chemical
TC2 | Speed vs capture certainty | #39 Productivity | #27 Reliability · #31 harmful side effects | 1, 35, 10, 38 · 35, 22, 18, 39 | 35 Parameter changes | speed profile: slow while the basket is full, faster as it drains, face velocity held at 5× | adopted
TC2 | Speed vs capture certainty | #39 Productivity | #27 Reliability · #31 harmful side effects | 1, 35, 10, 38 · 35, 22, 18, 39 | 10 Preliminary action | 300 rpm pre-spin seats the sleeve and the water before the ramp | adopted
TC2 | Speed vs capture certainty | #39 Productivity | #27 Reliability · #31 harmful side effects | 1, 35, 10, 38 · 35, 22, 18, 39 | 1 Segmentation | split the batch into small doses | rejected: pour already 10 s
TC2 | Speed vs capture certainty | #39 Productivity | #27 Reliability · #31 harmful side effects | 1, 35, 10, 38 · 35, 22, 18, 39 | 18 Mechanical vibration | silicone feet isolate the residual unbalance | kept
TC3 | Water speed vs simplicity | #9 Speed | #33 Ease of operation · #36 Device complexity | 32, 28, 13, 12 · 10, 28, 4, 34 | 10 Preliminary action | vanes welded into the basket; the sleeve slides past their tips | adopted
TC3 | Water speed vs simplicity | #9 Speed | #33 Ease of operation · #36 Device complexity | 32, 28, 13, 12 · 10, 28, 4, 34 | 13 · 34 | pleats in the sleeve act as the vanes and are thrown away with it | rejected by simulation
TC3 | Water speed vs simplicity | #9 Speed | #33 Ease of operation · #36 Device complexity | 32, 28, 13, 12 · 10, 28, 4, 34 | 28 Mechanics substitution | tangential jet spins the water up | rejected: pour energy too small
TC3 | Water speed vs simplicity | #9 Speed | #33 Ease of operation · #36 Device complexity | 32, 28, 13, 12 · 10, 28, 4, 34 | 32 Colour changes | dye tracer shows how far the water lags | bench test 1
TC4 | Sample size vs sleeve life | #26 Quantity of substance | #39 Productivity | 13, 29, 3, 27 | 3 Local quality | count three Ø12 mm coupons at marked places, not the whole sleeve | adopted
TC4 | Sample size vs sleeve life | #26 Quantity of substance | #39 Productivity | 13, 29, 3, 27 | 27 Cheap short-living objects | the sleeve is a consumable; track cost per litre | costed, slide 16
TC4 | Sample size vs sleeve life | #26 Quantity of substance | #39 Productivity | 13, 29, 3, 27 | 29 Pneumatics and hydraulics | back-flush the sleeve to extend its life | rejected: washes the sample off
TC4 | Sample size vs sleeve life | #26 Quantity of substance | #39 Productivity | 13, 29, 3, 27 | 13 The other way round | coarse pre-sock in the funnel for turbid water | later
TC5 | Filter area vs rotor mass | #5 Area of moving object | #1 Weight of moving object | 2, 17, 29, 4 | 17 Another dimension | pleated sleeve: 2.3× the area in the same basket | rejected by simulation
TC5 | Filter area vs rotor mass | #5 Area of moving object | #1 Weight of moving object | 2, 17, 29, 4 | 2 Taking out | keep the area, raise the head instead (speed profile) | adopted via TC2
TC6 | Plastic-free wet path vs complexity | #31 Object-generated harmful factors | #36 Device complexity | 19, 1, 31 | 1 Segmentation | bowl rolled from sheet with a welded floor, not deep-drawn | adopted for the prototype
TC6 | Plastic-free wet path vs complexity | #31 Object-generated harmful factors | #36 Device complexity | 19, 1, 31 | 31 Porous materials | perforated cage + woven mesh stay steel | kept
ID | Physical contradiction | Separation | Resolution in v1.1
PC1 | Pores must be small (catch 1 µm) and large (pass 15 L/h) | Space | two layers in series: each stays as open as 602 h, together they catch 98.9 % at 1–2 µm
PC2 | The basket must spin fast (flow) and slowly (paper rating, noise) | Time | speed profile: about 1,250 rpm when full, rising to 2,500 rpm as the water thins
PC3 | Particles must reach the paper and float inward (PP, PE > 20 µm) | Condition | slow, steady stop lets the floating layer settle onto the sleeve; count a floor coupon too
PC4 | The sleeve must be sealed and removable | Time + self-service (25) | spin pressure presses the 20 mm overlaps shut; floor gasket and clamp ring seal the ends; lifts out when stopped
Matrix cells from the TRIZ40 contradiction matrix (triz40.com/aff_Matrix_TRIZ.php), checked cell by cell.